# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title : Supply of Research equipment	p 1 /…

- Publication reference : Prior information notice 2015/S 213-388098

Column 1-2 should be completed by the Contracting Authority

Column 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 01 | Spark plasma sintering unit |  |  |  |
|  | Consisting of a fully functional spark plasma sintering furnace, to include double walled water cooled vessel, programmable heating system, hydraulic pressing system, vacuum gas system and integrated temperature measurement with the following technical characteristics: 1.1 Pressing force: minimum in the range of 2 … 25 kN 1.2 Piston stroke: minimum in the range of 0 … 50 mm 1.3 Maximum piston speed: at least 2 mm/s 1.4 Maximum mould dimensions: at least Ø 80 x 140 mm 1.5 Working temperature: at least up to 2000 °C 1.6 Maximum temperature: at least 2000 °C 1.7 Temperature measurement in axial direction: suitable for measuring at least in the range of 250 … 2500°C 1.8 Final vacuum in cold furnace: at least 5 x 10-2 mbar |  |  |  |
|  | 1.9 Working pressure: at least in the range of +20 - +60 mbar 1.10 Maximum cooling capacity: at least 15 kW 1.11 Maximum continuous DC current: at least 2.500 A 1.12 Maximum continuous power: at least 25 kW 1.13 Pulse duration (variable as a function of demanded heating capacity and the mould design): at least in the range of 1 ... 255 ms 1.14 Pause duration (variable as a function of demanded heating capacity and the mould design): at least in the range of 0 … 255 ms 1.15 Pulse number range (variable as a function of demanded heating capacity and the mould design): at least in the range of 1 … 255 |  |  |  |

| 02 | X ray diffraction unit |  |  |  |
| --- | --- | --- | --- | --- |
|  | Consisting of a fully functional diffractometer suitable for analyzing bulk metallic materials with the minimum dimensions in the range of 10x10x2 mm, with corresponding enclosure, goniometer, divergent beam optics to perform Bragg-Bretano measurements, tube housing and software for data collection with the following minimal requirements: 1.Enclosure to safely include the main components and accessories, the high-tension generator, high tension cables and mains supply, electronics and tube housing. 2. Goniometer with at least the following specifications: - accuracy +/- 0.0025 degrees or better, - reproducibility (without attachments) +/- 0.0001 degrees or better, - reproducibility (with attachments) +/- 0.001 degrees or better, - scan speed 0.001 to 1.27 degrees per second or better, - slew speed 12 degrees per second or better, - minimum step size 0.001 degrees or better, - 2theta range -40 to 220 degrees or better, - theta range -15 to 181 degrees or better, 3. PC and monitor with software to control the diffractometer hardware and to allow data collection with graphics features for examining, displaying and editing diffractograms 4. X-Ray tube with the maximum power at least 1.7 kW and filter for suppression of the unwanted radiation to less than 1% of the Cu K-alpha intensity - posibility to change from point focus to line focus without removal of the tube from its housing, without the need to dismount electrical or water feed, without the need for realignment; - HT generator of 3kW output or more; |  |  |  |
|  | 5. Incident beam path with holders and corresponding slits: - at least divergence slits of 4, 2, 1, 1/2 and 1/4 degrees and beam width masks of 20, 15, 10 and 5 mm - at least one soller slit of 0.04 radians 6. Divergent beam optic to perform Bragg-Brentano type measurements, monochromator (incident optical path): - energy resolution of 450 eV or better - to include corresponding fixed divergence slit holder, anti-scatter device, Soller slits holder, attenuator and a set of beam masks to reduce the beam width from X-ray mirror and hybrid monochromators. 7. Diffracted Beam Path: - at least one soller slit with holder, of 0.04 radians - at least antiscatter slits of 4, 2, 1, 1/2 and 1/4 degrees with holder - at least fixed receiving slit of 4, 2, 1, 1/2 and 1/4 degrees with holder, 8. At least detector 0D (including counting electronics); 9. Air cooled chiller for X-Ray Tube water circuit (closed loop); 10. Optical accessories (including X-Ray Tube) removal or replacement should not require realignment of the optical path to match the system original technical performance (auto calibration or fixed reference holders); 11. Installation, shipping and transport to Timisoara; 12. Warranty of at least 12 months (including spare parts, labour and transports costs). 13. Performance for instrument: resolution of 0.028° or better on a LaB6 (peak width) 14. X-Ray Safety: less than 1 μSv/h at 10 cm distance from the outside surface of the enclosure |  |  |  |
| 03 | High-vacuum microfabrication unit |  |  |  |
|  | Consisting of a fully functional arc melter with high vacuum chamber and pumping unit, rapid solidification by meltsuction capability, with the following minimal requirements: 1.1 High vacuum chamber with electrode holder consisting of: 1.1.1 Stainless steel chamber with movable electrode feedthrough and water-cooled base plate. 1.1.2 Volume: min 16 dm3 1.1.3 Total leak rate of the chamber: less than 1 x 10-6 mbar l/s. 1.1.4 Suitable for operation in inert gas (Ar, He, N) at least up to 2000 mbar. 1.1.5 Viewing ports: min. 1 of minimum 80 mm diameter, with removable eye protection glass 1.1.6 Copper plate with 3 different crucible shapes for the samples to be molten. 1.1.7 Manometer attached to the evacuation socket. 1.1.8 Flow control system for cooling of the chamber 1.1.9 Water-cooled, vacuum-tight and movable electrode holder, integrated in the top flange. 1.1.10 Electrical insulation to allow the safe ignition of the arc. 1.1.11 Electrode position adjustment: motor driven with min 50 mm displacement 1.1.12 Key switches for control of the motor drive, ignition and the generator current integrated in the operating handle 1.1.13 Possibility to manually change the electrode 1.1.14 Electrode material: W 1.1.15 Sample manipulator for operation without breaking the vacuum |  |  |  |
|  | 1.1.16 Suction casting capability to include at least: - vacuum-tight crucible mantle - cooling and lower cover - feedthrough for the evacuation of the crucible insert - evacuation tube. - device for lifting the crucible insert - within the crucible mantle. electric valve for suction casting. 1.1.17 Crucibles for rapid cooling, adapted for the suction casting component, with the following dimensions: - at least one with inner diameter: - 3 mm or smaller, - min length 150 mm - at least one with inner diameter: - 5 mm or smaller, - min length 150 mm 1.2. High vacuum System 1.2.1 Preliminary vacuum pump: -Nominal pumping speed: min 15 m3 h-1 - Pumping speed: min 12 m3 h-1 - Ultimate partial pressure without gas ballast: less than 1 x 10-4 mbar - Ultimate total pressure with gas ballast: less than 2 x 10-2 mbar 1.2.2 High vacuum pump: - Pumping speed N2: min 145 l sec-1 - Pumping speed He: min 135 l sec-1 - Ultimate pressure: less than 10-8 mbar 1.2.3 Measuring System Measuring range: at least from 5 x 10-9 mbar |  |  |  |
|  | 1.3. Arc Melting Generator 1.3.1 Configuration: to include HF ignition generator 1.3.2 Minimum setting range: at least up to 350 A |  |  |  |

Unless otherwise specified, the requirements in this Technical Specifications are presented as minimum standard which the offered goods must meet. The offers that will not satisfy completely and correctly the minimum requirements mentioned bellow will be disqualified.
	When brands names/specifications referring to or describing products of a given brand or origin are used in this technical specifications, they are used in descriptive purposes only and should be interpreted as equivalent.